Corrosion test panel rating record sheet
Instructions
Use one sheet per exposure run and inspection. Rate the control panel first.
Rate rust (ASTM D610), blistering (ASTM D714) and scribe creep (ASTM D1654 or ISO 4628-8) against the standards' own pictorial references, kept beside the panel. This sheet does not reproduce them. Record the scribe procedure and the method used to remove loose coating before measuring creep.
Write each rating exactly as the standard designates it. Keep ASTM and ISO 4628 ratings apart.
Fill section 2 before exposure starts. Do not change a criterion after panels are rated.
Leave a rating blank and give the reason in Notes rather than guess. Transfer entries to the Excel log the same day.
The example page at the end is illustrative only and is kept apart from the blank form.
1. Run and inspection
Exposure test and edition: __________________________________________________
Cabinet or site ID: __________________________________________________
Exposure start date (YYYY-MM-DD): __________________________________________________
Exposure hours at this inspection (h): __________________________________________________
Inspection date (YYYY-MM-DD): __________________________________________________
Rater (name): __________________________________________________
Second rater, if used (name): __________________________________________________
Lighting and viewing conditions: __________________________________________________
Substrate and pretreatment: __________________________________________________
Scribe type: __________________________________________________
Scribe method, procedure and removal: __________________________________________________
2. Acceptance criteria (as frozen before exposure)
Rust criterion (ASTM D610 grade or ISO 4628-3): __________________________________________________
Blister criterion (ASTM D714 size and density or ISO 4628-2): __________________________________________________
Scribe creep criterion (mm, mean or maximum): __________________________________________________
Date criteria were frozen (YYYY-MM-DD): __________________________________________________
Agreed by (both parties): __________________________________________________


3. Panel ratings
Rust distribution: S spot, G general, P pinpoint, H hybrid (ASTM D610). Judgment: met, borderline, not met, unassessed.
	Panel ID
	Replicate group
	Control? [ ]
	System
	DFT (µm)
	Rust grade D610 (0-10)
	Rust distr. [ ]S [ ]G [ ]P [ ]H
	Blister size D714
	Blister density D714
	Creep mean (mm)
	Creep max (mm)
	Scribe rating D1654 (0-10)
	Photo ref.
	Judgment [ ]met [ ]bord. [ ]not met [ ]unass.
	Notes

	
	
	
	
	
	
	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	
	
	
	
	
	
	



4. Run check and sign-off
[ ] Control panel rated first and within its known range
[ ] Cabinet records checked for deviations during this interval
[ ] Every panel photographed under the same light
[ ] Edges and hanger holes noted, not rated
[ ] Entries transferred to the Excel log
Signed (name, date): __________________________________________________


Example (illustrative only)
Every value below is invented to show how the sheet reads. None comes from a real test.
	Panel ID
	Replicate group
	Control
	System
	DFT (µm)
	Rust grade
	Rust distr.
	Blister size
	Blister density
	Creep mean (mm)
	Creep max (mm)
	Scribe rating
	Photo ref.
	Judgment
	Notes

	P-001
	Control
	yes
	Reference R
	175
	9
	S
	10
	
	1.0
	1.6
	8
	P-001_1000h.jpg
	met
	Rated first

	P-014
	SYS-B rep 2 of 3
	no
	SYS-B
	180
	9
	S
	8
	F
	1.2
	2.0
	7
	P-014_1000h.jpg
	met
	Hanger hole rust, not rated

	P-015
	SYS-B rep 3 of 3
	no
	SYS-B
	165
	8
	P
	8
	M
	2.0
	3.4
	7
	P-015_1000h.jpg
	borderline
	At the limit; second rating under the same light asked



Illustrative run: ASTM B117-26a, cabinet 2, 1000 h, criterion creep mean at most 2 mm. Substrate cold-rolled steel, zinc phosphate; single diagonal scribe to metal; scribe method ASTM D1654-24e1, removal per lab procedure LP-07.
Template from www.lubinpla.com/post/corrosion-test-panel-rating-record-sheet




